PULSE-SAFE-001

Clinical Safety Management, Hazard Analysis & Safety Case

Pulse Health Platform

| Document Control | Value |
| --- | --- |
| Parent documents | Pulse Master Blueprint; PULSE-ARCH-001; PULSE-DATA-001; PULSE-CLIN-001 |
| Version | 1.0 |
| Status | Clinical Safety Engineering Baseline |
| Audience | Clinical Safety, Clinical Leadership, Engineering, Product, QA, Security, Operations |
| Purpose | Define the safety management system, hazard model, controls, evidence and release gates |
| Authority | No feature may be released when an unacceptable clinical safety risk remains without approved disposition |

# 1. Purpose, Scope and Safety Position

Pulse is safety-related healthcare software. A defect, ambiguous workflow, delayed message, stale record, incorrect identity, misleading visualization, failed interface or poorly designed human interaction can contribute to patient harm even when Pulse does not directly deliver treatment. Clinical safety therefore operates as a lifecycle engineering discipline, not as a final QA checklist.

This specification establishes the Clinical Safety Management System (CSMS) for Pulse: governance, hazard identification, risk estimation, controls, verification, residual-risk acceptance, incident learning and the safety case required for release and deployment.

SAFETY PRINCIPLE — Safety claims must be supported by evidence. Absence of reported incidents is not evidence that a hazard is controlled.

# 2. Safety Objectives

Prevent or reduce foreseeable patient harm caused or contributed to by Pulse.

Make clinical responsibility, patient identity, status and provenance unambiguous.

Fail safely when dependencies, networks, devices or integrations fail.

Ensure critical information is delivered through accountable workflows rather than passive display alone.

Preserve evidence when clinical data is corrected, amended or exchanged.

Design controls around real clinical work and human limitations.

Detect hazardous degradation before it becomes silent failure.

Provide a defensible safety argument for each production release and deployment.

# 3. Safety Governance

| Role | Primary responsibility |
| --- | --- |
| Clinical Safety Officer / equivalent | Owns clinical safety process, hazard log and safety case; independent authority to escalate unacceptable risk. |
| Chief Medical/Clinical representative | Clinical validity, workflow/protocol approval and residual-risk acceptance. |
| Nursing safety representative | Nursing workflow, eMAR, handover, deterioration and bedside safety. |
| Pharmacy safety representative | Medication, formulary, dispensing and administration risks. |
| Laboratory/Radiology representatives | Diagnostic workflow, validation and critical-result risks. |
| Product owner | Ensures safety requirements remain in scope and traceable. |
| Engineering lead | Implements technical controls and supplies verification evidence. |
| QA/Validation lead | Independent verification, regression and safety test evidence. |
| Security/Privacy leads | Cyber/privacy hazards that can affect clinical safety. |
| Facility deployment lead | Local configuration, readiness, training and go-live risk. |
| Incident review group | Post-market event review, corrective action and learning. |

## 3.1 Independence and escalation

A safety-critical defect cannot be waived solely by delivery pressure.

Residual risk acceptance requires the designated clinical/governance authority, not only engineering.

The person implementing a safety control should not be the only person validating it.

Safety concerns may trigger release stop, feature disablement, rollback or restricted deployment.

# 4. Safety Lifecycle

Clinical need / change request

↓

Safety impact screening

↓

Hazard identification

↓

Risk estimation

↓

Control design

↓

Implementation

↓

Verification & validation

↓

Residual risk evaluation

↓

Safety case / release decision

↓

Deployment monitoring

↓

Incident / feedback / change

↺ back to hazard analysis

Every material workflow, architecture, data-model, interface or configuration change is screened for clinical safety impact. The hazard log is a living product artefact.

# 5. Hazard Taxonomy

| Category | Examples |
| --- | --- |
| Identity | Wrong patient selected, duplicate identity, wrong MRN, incorrect merge. |
| Clinical information | Missing, stale, misleading, incorrectly amended or wrongly attributed information. |
| Medication | Allergy miss, wrong dose/unit/route, duplicate therapy, wrong patient, missed administration. |
| Diagnostics | Wrong specimen, wrong-patient result, delayed result, critical result not acknowledged. |
| Imaging | Wrong study/patient, wrong laterality, unavailable images, critical finding not escalated. |
| Referral/transfer | Lost referral, false acceptance, inadequate transfer package, responsibility gap. |
| ADT/bed | Wrong location, missing handover, bed assignment mismatch, discharge error. |
| Procedure/theatre | Wrong site/laterality, missing checklist, implant/batch traceability failure. |
| Anatomy visualization | Wrong structure/laterality, misleading geometry, stale model mapping. |
| Decision support | False positive fatigue, false negative, opaque recommendation, stale rule. |
| Interoperability | Message loss, duplication, semantic mismatch, unit/code transformation error. |
| Offline/synchronisation | Stale record, conflicting edits, duplicate order, missing event after reconnect. |
| Human factors | Poor salience, ambiguous controls, excessive clicks, copy-forward, alert fatigue. |
| Security availability | Ransomware/outage prevents access to critical information. |
| Privacy-related safety | Over-restriction hides necessary data; inappropriate disclosure changes care behaviour. |
| Operations/inventory | False stock availability, expired/recalled product, reagent shortage not visible. |
| Configuration | Unsafe local template, wrong reference range, wrong escalation timer, incorrect privilege. |

# 6. Risk Estimation Model

## 6.1 Severity scale

| Level | Definition | Illustrative consequence |
| --- | --- | --- |
| S1 Negligible | No injury or clinically insignificant inconvenience. | Minor delay without effect on care. |
| S2 Minor | Temporary minor harm or additional low-risk intervention. | Repeat specimen or short treatment delay. |
| S3 Moderate | Temporary harm requiring significant intervention or prolonged care. | Missed dose causing deterioration; delayed diagnosis. |
| S4 Major | Serious permanent harm, life-threatening deterioration or major intervention. | Wrong-site intervention narrowly/actually causing severe harm. |
| S5 Catastrophic | Death or multiple severe harms. | Critical transfer/result failure contributing to death. |

## 6.2 Probability scale

| Level | Definition |
| --- | --- |
| P1 Rare | Exceptional; credible but not expected in normal lifecycle. |
| P2 Unlikely | Could occur but infrequent with existing controls. |
| P3 Possible | May occur periodically under realistic conditions. |
| P4 Likely | Expected to occur repeatedly unless controlled. |
| P5 Frequent | Occurs commonly or is inherent in workflow without control. |

## 6.3 Risk matrix

| Probability \ Severity | S1 | S2 | S3 | S4 | S5 |
| --- | --- | --- | --- | --- | --- |
| P5 Frequent | M | H | H | E | E |
| P4 Likely | M | M | H | E | E |
| P3 Possible | L | M | H | H | E |
| P2 Unlikely | L | L | M | H | H |
| P1 Rare | L | L | M | M | H |

L = Low, M = Moderate, H = High, E = Extreme. Product governance may refine numerical scoring, but the qualitative meaning and escalation must remain consistent.

## 6.4 Risk acceptance

| Residual level | Disposition |
| --- | --- |
| Low | May be accepted by designated product/clinical safety authority with evidence. |
| Moderate | Requires documented ALARP-style rationale/benefit, control verification and clinical approval. |
| High | Not releasable without further risk reduction or formally exceptional executive/clinical governance decision where legally permissible. |
| Extreme | Unacceptable for production use; release/deployment blocked. |

# 7. Hazard Record Schema

| Field | Required content |
| --- | --- |
| Hazard ID | Stable identifier, e.g., HAZ-ID-001. |
| Hazard | Potential source/situation of harm. |
| Initiating causes | Technical, human, workflow, configuration or external causes. |
| Hazardous situation | How patient/user becomes exposed. |
| Potential harm | Clinical consequence. |
| Affected workflow | Encounter/order/referral/etc. |
| Initial severity/probability | Pre-control assessment. |
| Controls | Preventive, detective, mitigative controls. |
| Safety requirements | Traceable SHALL statements. |
| Verification evidence | Tests, reviews, simulations, monitoring. |
| Residual severity/probability | Post-control assessment. |
| Residual risk | L/M/H/E. |
| Owner | Named accountable role. |
| Status | Open/controlled/accepted/retired. |
| Linked changes/incidents | Requirements, tickets, incidents, CAPA. |

# 8. Safety Control Hierarchy

Eliminate the hazardous condition by design where possible.

Prevent the hazardous action/state using constraints and strong defaults.

Detect the hazardous state early and route accountable action.

Mitigate consequence through escalation, redundancy or recovery.

Warn the user only when stronger controls are impractical or inappropriate.

Use training/procedure as a supporting control, not the sole control for predictable software hazards.

SAFETY PRINCIPLE — A warning message is not an adequate substitute for a design that can prevent a predictable catastrophic error.

# 9. Master Hazard Register - Identity and Context

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-ID-001 | Wrong patient selected due to similar names/search results | Care documented/administered to wrong patient | Persistent patient banner; multi-factor identity confirmation; constrained search; bedside verification | M/L |
| HAZ-ID-002 | Duplicate patient created | Fragmented history; allergy/results missed | Search-before-create; duplicate detection; provisional identity; merge governance | M |
| HAZ-ID-003 | Incorrect patient merge | Records combined across people | Restricted merge privilege; side-by-side review; reversible link history; audit | L/M |
| HAZ-ID-004 | Stale patient context remains after navigation | Order/note entered for previous patient | Context reset; prominent banner; confirmation on high-risk actions; automated UI tests | L/M |
| HAZ-ID-005 | Provisional emergency identity never reconciled | Fragmented care and duplicate medication/results | Visible provisional status; reconciliation work queue; discharge gate warning | M |

# 10. Master Hazard Register - Clinical Documentation

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-CLN-001 | Signed note overwritten | Loss of clinical history/accountability | Immutable signed version; addendum/amendment workflow | L |
| HAZ-CLN-002 | Copy-forward propagates outdated/incorrect data | Wrong clinical decisions | Source/time labels; limit copy-forward; clinician confirmation; audit | M |
| HAZ-CLN-003 | Effective time confused with entry time | Incorrect sequence of events | Separate effective_at/recorded_at; late-entry labels | L/M |
| HAZ-CLN-004 | Important allergy/problem hidden in long record | Medication/treatment error | Patient banner; prioritised safety summary; role-specific views | M |
| HAZ-CLN-005 | Draft content mistaken as final | Premature clinical action | Strong draft/final visual state; API status; downstream rules | L |

# 11. Master Hazard Register - Medication

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-MED-001 | Known allergy not considered | Severe allergic reaction | Structured allergy; prescribing alert; pharmacist verification; override audit | M |
| HAZ-MED-002 | Dose/unit error | Overdose/underdose | Structured dose/unit; range checks where validated; pharmacist review; high-risk medication rules | M |
| HAZ-MED-003 | Wrong route/frequency | Treatment failure/harm | Structured route/frequency; order validation; eMAR display | M |
| HAZ-MED-004 | Duplicate therapy | Adverse effect | Medication reconciliation; duplication detection; pharmacist intervention | M |
| HAZ-MED-005 | Dispense interpreted as administration | False treatment record | Separate Dispense and Administration entities/workflows | L |
| HAZ-MED-006 | Administration recorded for wrong patient | Direct medication harm | Bedside identity check; patient/medication context; scan support where deployed | M |
| HAZ-MED-007 | Stopped order remains on eMAR | Unintended dose | Event-driven schedule update; stopped state prominence; reconciliation | M |
| HAZ-MED-008 | Unsafe substitution during stockout | Therapeutic error | Substitution policy; pharmacist/prescriber approval; product equivalence rules | M |

# 12. Master Hazard Register - Laboratory and Diagnostics

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-LAB-001 | Specimen labelled for wrong patient | Wrong diagnosis/treatment | Bedside label workflow; identifiers; accession checks; rejection process | M |
| HAZ-LAB-002 | Result attached to wrong patient/order | Wrong treatment | Interface identity reconciliation; accession/order constraints; audit | M |
| HAZ-LAB-003 | Critical result not acknowledged | Delayed life-saving intervention | Critical flag + accountable task + escalation timers + audit | M |
| HAZ-LAB-004 | Final result silently edited | Clinician acts on untraceable data | Amended/corrected result version; notification to prior viewers where appropriate | L/M |
| HAZ-LAB-005 | Unit/reference-range mismatch | Misinterpretation | Coded units; test catalogue; analyzer mapping validation; reference context | M |
| HAZ-LAB-006 | Analyzer/interface outage creates false completeness | Missed tests | Interface health; pending queues; downtime status; reconciliation | M |

# 13. Master Hazard Register - Imaging

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-IMG-001 | Wrong-patient imaging association | Misdiagnosis/intervention | DICOM identity reconciliation; accession/worklist matching; exception queue | M |
| HAZ-IMG-002 | Wrong laterality/site in order/report | Wrong-site treatment | Structured laterality; display throughout workflow; report verification | M |
| HAZ-IMG-003 | Images unavailable but UI implies available | Decision made without images | Explicit availability state; PACS health; viewer failure handling | L/M |
| HAZ-IMG-004 | Critical imaging finding not escalated | Delayed treatment | Critical finding workflow equivalent to critical lab result | M |

# 14. Master Hazard Register - Referral and Transfer

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-REF-001 | Referral lost/not delivered | Delay in specialist care | Durable case state; delivery acknowledgement; queue monitoring; retry/fallback | M |
| HAZ-REF-002 | Acceptance mistaken for care assumption | Responsibility gap | Explicit responsibility milestones; arrival/care-assumed state | L/M |
| HAZ-REF-003 | Incomplete referral package | Receiver unprepared/unsafe transfer | Required clinical summary; package review; more-info state | M |
| HAZ-REF-004 | Stale capacity/bed information | Unsafe or failed transfer | Time-stamped capacity; receiving confirmation/reservation | M |
| HAZ-REF-005 | Patient deteriorates in transit without adequate plan | Severe harm | Transport requirements; escort/equipment; stabilization; ETA communication | M/H |
| HAZ-REF-006 | Sensitive data disclosed beyond need | Privacy harm affecting care/trust | Case-scoped package; consent/legal basis; manifest; audit | M |
| HAZ-REF-007 | Outcome not returned to referring team | Broken continuity | Outcome task; closure criteria; reminders/escalation | M |

# 15. Master Hazard Register - Admission, Handover and Discharge

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-ADT-001 | Bed assignment treated as responsibility transfer | Patient unattended | Ward arrival + handover acknowledgement required | L/M |
| HAZ-ADT-002 | Pending tasks omitted from handover | Missed medication/result/action | Generated pending-work list + manual review + acknowledgement | M |
| HAZ-ADT-003 | Discharge closes pending critical result | Missed diagnosis/treatment | Result-review queue independent of encounter closure | M |
| HAZ-ADT-004 | Medication reconciliation incomplete at discharge | Medication error/readmission | Reconciliation workflow and discharge medication list | M |
| HAZ-ADT-005 | Financial status blocks truthful clinical discharge | Incorrect census/bed state | Separate clinical and financial lifecycle states | L |
| HAZ-ADT-006 | Wrong bed/location shown | Delayed response/wrong patient contact | Single active placement rule; transfer transaction; reconciliation | L/M |

# 16. Master Hazard Register - Theatre and Procedures

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-THR-001 | Wrong patient/site/laterality | Catastrophic procedure error | Identity/site verification; structured laterality; checklist; anatomy linkage | M |
| HAZ-THR-002 | Checklist auto-completed or bypassed | Safety step falsely documented | Human-attributable completion; no auto-tick; exception audit | L/M |
| HAZ-THR-003 | Implant batch not traceable | Recall/follow-up failure | Batch/lot consumption linked to patient/procedure | L |
| HAZ-THR-004 | Specimen lost/misidentified | Missed/wrong pathology diagnosis | Specimen chain-of-custody and label verification | M |
| HAZ-THR-005 | Recovery handover incomplete | Post-op deterioration | Structured recovery criteria + handover acknowledgement | M |

# 17. Master Hazard Register - Anatomical Mapping

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-ANAT-001 | Wrong anatomical structure selected | Misleading diagnosis/procedure plan | Semantic structure label; hierarchy; confirmation; note context | M |
| HAZ-ANAT-002 | Wrong laterality | Wrong-side intervention/documentation | Explicit laterality; paired-structure validation; theatre cross-check | M |
| HAZ-ANAT-003 | 3D coordinate interpreted as exact clinical truth | False precision | Store semantic structure + narrative; geometry labelled illustrative/documentary | L/M |
| HAZ-ANAT-004 | Model version changes invalidate coordinates | Historical finding misplaced | Persist model/version/structure key; migration rules; preserve old viewer assets | L |
| HAZ-ANAT-005 | Visual model used as autonomous diagnosis | Clinical misjudgment | Product wording; no autonomous diagnosis; clinician accountable | L/M |

# 18. Master Hazard Register - Decision Support and AI

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-CDS-001 | False negative recommendation | Missed condition | CDS never substitutes required clinical assessment; rule validation/monitoring | M |
| HAZ-CDS-002 | False positive overload/alert fatigue | Important alert ignored | Severity tiers; tuning; override analysis; alert burden metrics | M |
| HAZ-CDS-003 | Stale guideline/rule | Outdated care recommendation | Versioned rule source; owner/review date; expiry/disable control | M |
| HAZ-AI-001 | Generated summary omits/misstates fact | Clinical error | Source-linked review; label generated content; clinician confirmation | M |
| HAZ-AI-002 | AI suggestion written as fact without approval | False diagnosis/plan | No silent write to signed record; explicit acceptance/edit | L/M |
| HAZ-AI-003 | Automation bias | Clinician over-relies on system | Human factors, uncertainty display, training, audit/monitoring | M |

# 19. Master Hazard Register - Offline, Integration and Infrastructure

| ID | Hazard / causes | Potential harm | Key controls | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-OFF-001 | Stale offline record used | Medication/decision error | Last-sync indicator; freshness rules; restricted high-risk actions when stale | M |
| HAZ-OFF-002 | Conflicting offline edits | Lost/contradictory clinical data | Conflict policy; append semantics; human reconciliation for clinical conflict | M |
| HAZ-OFF-003 | Duplicate order after reconnect | Duplicate test/medication | Idempotency keys; local IDs; reconciliation | L/M |
| HAZ-INT-001 | Message delivered twice | Duplicate action | Idempotent consumers; event IDs; deduplication | L |
| HAZ-INT-002 | Semantic mapping error | Wrong unit/code/result | Versioned mappings; conformance tests; quarantine exceptions | M |
| HAZ-INF-001 | Platform outage removes critical record access | Delayed/unsafe care | HA/edge/downtime procedure; backups; recovery objectives | M |
| HAZ-INF-002 | Clock/timezone mismatch | Wrong sequence/dose schedule | NTP/time health; UTC storage; local display; drift monitoring | L/M |

# 20. Human Factors Safety

Critical patient identity, allergies, encounter context and responsibility remain visible at decision points.

Do not rely on colour alone for severity/status.

Destructive/high-risk actions use consequence-specific confirmation, not generic 'Are you sure?'.

Avoid modal/alert saturation; prioritise clinically meaningful interruptions.

Keyboard and scanner workflows are supported where speed reduces workaround risk.

Auto-save must not imply signature/finalisation.

Defaults must be clinically safe and visible; never preselect dangerous values merely for speed.

Copy/paste and copy-forward are constrained and source-labelled.

Long lists must support search/filter while preserving high-risk flags.

Training environments use synthetic data and realistic workflow simulations.

# 21. Safety Requirements Catalogue

| ID | Requirement |
| --- | --- |
| SR-001 | Pulse SHALL maintain persistent patient identity context during every patient-specific clinical action. |
| SR-002 | Pulse SHALL preserve signed/final clinical history through amendment rather than destructive overwrite. |
| SR-003 | Pulse SHALL maintain a critical-result acknowledgement and escalation mechanism independent of notification delivery. |
| SR-004 | Pulse SHALL distinguish medication prescribing, dispensing and administration as separate attributable events. |
| SR-005 | Pulse SHALL prevent stock ledger entries from being destructively edited after posting. |
| SR-006 | Pulse SHALL represent referral acceptance, physical arrival and assumption of care as distinct states. |
| SR-007 | Pulse SHALL preserve source organisation and provenance for externally received clinical information. |
| SR-008 | Pulse SHALL visibly indicate degraded/offline/stale-data state. |
| SR-009 | Pulse SHALL require explicit laterality where clinically applicable to paired anatomical structures/procedures. |
| SR-010 | Pulse SHALL prevent a signed theatre safety checklist from being auto-completed by system inference. |
| SR-011 | Pulse SHALL maintain pending-result review responsibility after encounter/discharge closure. |
| SR-012 | Pulse SHALL retain audit evidence for break-glass access and cross-facility disclosure. |
| SR-013 | Pulse SHALL use idempotent processing for retried safety-relevant integration messages. |
| SR-014 | Pulse SHALL preserve effective clinical time separately from system recording time. |
| SR-015 | Pulse SHALL not convert AI-generated content into signed clinical fact without authorised human confirmation. |

# 22. Verification Evidence Model

| Control type | Evidence examples |
| --- | --- |
| Data constraint | Schema tests, migration tests, property tests. |
| State machine | Transition unit/integration tests; invalid transition tests. |
| UI safety | Usability test, screenshot review, simulated clinical scenario. |
| Alert/escalation | Timer simulation, delivery failure test, acknowledgement audit. |
| Integration | Contract/conformance tests, replay/duplicate tests, mapping validation. |
| Offline | Network-loss simulations, conflict tests, reconnect reconciliation. |
| Security | Threat tests, privilege tests, break-glass audit. |
| Performance | Load/latency tests for safety-critical actions. |
| Operational | Backup restore, failover, downtime drill. |
| Clinical | Clinical review, scenario validation, protocol sign-off. |

# 23. Safety Traceability

Hazard

↓

Safety Requirement (SR-xxx)

↓

Architecture / Data / Workflow control

↓

Implementation requirement / ticket

↓

Verification test

↓

Evidence artefact

↓

Residual risk

↓

Safety case claim

Every High/Extreme initial-risk hazard requires explicit traceability through this chain. A release cannot rely on an undocumented assumption as a control.

# 24. Clinical Safety Case Structure

| Section | Required argument/evidence |
| --- | --- |
| System definition | What version/configuration/deployment is being assessed. |
| Intended use | Users, care settings, capabilities and exclusions. |
| Safety governance | Responsible roles and approvals. |
| Hazard analysis | Current hazard log and methodology. |
| Risk controls | Implemented preventive/detective/mitigative controls. |
| Verification | Evidence that controls work. |
| Known limitations | Residual limitations, unsupported workflows, downtime constraints. |
| Residual risk | Accepted risks and rationale. |
| Operational readiness | Training, support, monitoring, incident response. |
| Change assessment | Differences from prior approved release. |
| Conclusion | Signed statement whether system is acceptably safe for intended deployment. |

# 25. Safety Release Gates

| Gate | Minimum requirement |
| --- | --- |
| Design complete | Safety impact assessment and hazards identified for material features. |
| Code complete | Safety requirements implemented; no known Extreme residual risk. |
| Validation ready | Traceability and test cases available; clinical scenarios defined. |
| Release candidate | Safety tests passed; open High risks dispositioned; regression complete. |
| Production approval | Safety case approved; deployment configuration reviewed; rollback/downtime ready. |
| Post-go-live | Enhanced monitoring window active; incident escalation contacts confirmed. |

SAFETY PRINCIPLE — A passing functional test suite is necessary but not sufficient for clinical safety release.

# 26. Change Control and Safety Impact Assessment

| Change type | Safety review expectation |
| --- | --- |
| UI wording/layout | Assess misinterpretation, salience and wrong-patient/context risk. |
| Clinical workflow/state | Full hazard review and scenario regression. |
| Data model | Assess provenance, migration, history and interoperability. |
| Terminology/reference range | Clinical approval and versioned deployment. |
| Medication/CDS rule | Clinical/pharmacy approval plus targeted safety validation. |
| Integration mapping | Conformance + semantic mapping validation. |
| Offline/sync logic | Conflict/duplicate/stale-data hazard testing. |
| Infrastructure | Availability, latency, recovery and data-loss impact. |
| AI model/prompt | Clinical performance, failure modes, provenance and human oversight. |
| Facility configuration | Local safety validation before activation. |

# 27. Incident Management

Safety event / near miss reported

↓

Triage severity + immediate containment

↓

Preserve logs/data/configuration/version evidence

↓

Clinical + technical investigation

↓

Identify contributing factors / root causes

↓

Corrective & Preventive Action (CAPA)

↓

Verify fix

↓

Update hazard log / safety case / training

↓

Communicate affected deployments where required

Near misses are safety data and should be reviewed, not discarded because no harm occurred.

Incident review avoids blame-oriented design and examines system, workflow, configuration, training and organisational contributors.

Potentially affected patient cohorts must be queryable when a defect may have influenced prior care.

Emergency containment may include feature flag disablement, workflow restriction, rollback or advisory.

# 28. Safety Monitoring Metrics

| Metric | Safety intent |
| --- | --- |
| Wrong-patient corrections | Detect identity/context weakness. |
| Duplicate patient rate | Identity quality. |
| Critical-result acknowledgement time | Escalation performance. |
| Unacknowledged critical results | Direct safety backlog. |
| Medication alert override rate | Alert quality/fatigue. |
| Stopped-order administrations | eMAR/order propagation risk. |
| Referral delivery/acceptance latency | Continuity risk. |
| Referrals stuck by state | Lost workflow detection. |
| Unreviewed results after discharge | Follow-up risk. |
| Offline conflict count | Sync safety. |
| Interface quarantine/error rate | Semantic/integration risk. |
| Handover acknowledgement gaps | Responsibility risk. |
| Break-glass use | Access/safety governance. |
| High-severity incident/near miss trend | Overall safety signal. |

# 29. Deployment Safety Readiness

Facility workflows mapped against PULSE-CLIN-001.

Roles/privileges validated using real job functions.

Local order/test/service catalogues clinically reviewed.

Medication formulary and units reviewed.

Critical-result thresholds and escalation contacts configured.

Reference ranges validated by laboratory leadership.

Referral partners/destinations verified.

Downtime procedures rehearsed.

Patient identity/merge policy approved.

Training completed for high-risk roles.

Go-live support and clinical escalation rota active.

Rollback/feature-disable mechanisms tested.

# 30. Safety in Multi-Facility and Network Operation

Source organisation is always visible for external data.

Receiving clinicians can distinguish local verified data from imported/referral information.

Cross-facility patient matching does not automatically merge legal records.

Referral package versions preserve exactly what the receiver saw.

Consent/disclosure evidence is linked to the exchange.

Network outage produces explicit undelivered/queued status.

A hospital cannot silently modify another hospital's signed clinical record.

Outcome information returned to the sender carries receiving provenance.

# 31. Safety in Analytics and Research

Analytics does not write back clinical conclusions without governed workflow.

Cohort/metric definitions are versioned.

Research datasets are separated from production clinical decision pathways unless explicitly approved.

Predictive models are evaluated for calibration, subgroup performance and drift appropriate to intended use.

A dashboard anomaly is not a clinical diagnosis; patient-level action requires access to underlying clinical context.

# 32. Safety in Patient-Facing Features

Patient portal clearly distinguishes clinician-authored facts from educational/generated content.

Critical patient instructions are not dependent solely on push notification delivery.

Proxy access is explicit and time/relationship governed.

Result-release policies prevent ambiguous provisional/final status.

Patients can report demographic/record concerns without directly overwriting signed records.

Emergency symptoms are not routed through asynchronous messaging as a substitute for emergency care instructions.

# 33. Safety Assurance for AI-Augmented Features

| Area | Required assurance |
| --- | --- |
| Summarisation | Source traceability; omission/hallucination evaluation; clinician review. |
| Coding assistance | Suggestion only until accepted; terminology/version displayed. |
| Risk prediction | Intended population, calibration, thresholds, drift monitoring, override. |
| Inventory forecasting | No direct clinical substitution; human procurement control. |
| Documentation assistance | No autonomous signature; authorship remains clinician. |
| Referral assistance | Cannot invent clinical facts or acceptance/capacity. |
| Anatomy assistance | Cannot convert visual inference to diagnosis without clinician action. |

SAFETY PRINCIPLE — Generative output must never be presented as if it were an observed clinical fact unless an authorised clinician has reviewed and adopted it.

# 34. Example Detailed Hazard Analysis

| Field | Example: Critical laboratory result not acted upon |
| --- | --- |
| Hazard ID | HAZ-LAB-003 |
| Initiating causes | Result released after clinician leaves; push fails; queue hidden; responsibility unclear. |
| Hazardous situation | Critical potassium result exists but no accountable clinician reviews it. |
| Potential harm | Arrhythmia, deterioration, death. |
| Initial risk | S5 / P3 = Extreme. |
| Preventive controls | Validated critical threshold; responsible team resolution; result-review queue. |
| Detective controls | Critical-result task; acknowledgement timer; dashboard; monitoring. |
| Mitigative controls | Escalation to backup clinician/ward/consultant; alternative contact path. |
| Safety requirements | SR-003, SR-011. |
| Verification | Threshold tests; notification-failure simulation; escalation timer test; discharge scenario. |
| Residual target | S5 / P1 = High; requires clinical governance acceptance because severity cannot be reduced, only probability. |
| Operational monitoring | Acknowledgement time, overdue critical results, escalation failures. |

# 35. Example Detailed Hazard Analysis - Referral

| Field | Example: accepted referral but patient has no receiving team on arrival |
| --- | --- |
| Hazard ID | HAZ-REF-002 |
| Initiating causes | Acceptance treated as transfer of responsibility; shift change; bed changed; ETA not updated. |
| Hazardous situation | Patient travels while sender assumes receiver owns care and receiver has not assumed care. |
| Potential harm | Treatment delay, deterioration, abandonment. |
| Initial risk | S4/S5, P2-P3 = High/Extreme. |
| Controls | Distinct accepted/departed/arrived/care-assumed states; receiving assignment; ETA; handover acknowledgement; escalation. |
| Verification | Scenario simulation across shift change/network delay; state-transition tests. |
| Residual target | Moderate or lower depending deployment. |
| Monitoring | Arrival-to-care-assumed time; referrals stuck in transit; cancelled resource reservations. |

# 36. Example Detailed Hazard Analysis - Anatomy

| Field | Example: wrong laterality from anatomical viewer |
| --- | --- |
| Hazard ID | HAZ-ANAT-002 |
| Initiating causes | Mirrored model orientation; user selects right structure while viewing patient's left; laterality hidden. |
| Hazardous situation | Wrong-side finding propagates to procedure planning. |
| Potential harm | Wrong-site procedure. |
| Initial risk | S5 / P2 = High. |
| Controls | Explicit textual laterality; orientation labels; structure metadata; procedure-site verification; checklist cross-check. |
| Verification | Usability testing; mirrored-view tests; left/right automated tests; theatre scenario. |
| Residual target | Moderate. |
| Monitoring | Laterality corrections; wrong-site near misses. |

# 37. Safety Requirements for Offline Mode

Show last successful synchronisation time and connectivity state.

Define which actions remain allowed when required data is stale.

Preserve local event IDs/idempotency keys.

Do not auto-resolve clinically meaningful conflicts by last-write-wins.

Reconcile medication orders/admins, results, referrals and admissions using domain-specific rules.

Queue outbound messages with visible delivery state.

Escalate records that fail reconciliation.

Retain downtime provenance on later-entered records.

# 38. Safety Requirements for Interoperability

Validate sender identity and message schema.

Validate patient and encounter context.

Preserve source code/unit/display before transformation.

Version terminology mappings.

Quarantine semantically invalid messages rather than coercing silently.

Use idempotency/deduplication.

Expose interface health and backlog.

Reconcile acknowledgements for safety-relevant exchanges.

Test with realistic partner data before production activation.

# 39. Safety Requirements for Configuration

Configuration has draft/review/approved/published states.

Safety-relevant configuration requires named clinical approver.

Effective dates and version are retained.

Reference range, medication, escalation and workflow changes are regression tested.

Production configuration export is reviewable before go-live.

Emergency rollback to prior configuration is supported where feasible.

Facility administrators cannot redefine platform safety invariants.

# 40. Clinical Safety Testing Catalogue

| Test family | Examples |
| --- | --- |
| Identity | Similar names, twins, duplicate MRNs, provisional patient, merge reversal. |
| Medication | Allergy, decimal dose, unit mismatch, stopped order, partial dispense, refusal. |
| Lab | Rejected specimen, critical result, corrected result, analyzer outage. |
| Imaging | Wrong accession, PACS outage, critical report, laterality. |
| Referral | No acknowledgement, more-info loop, decline, network loss, ETA change, arrival. |
| ADT | No bed, ward transfer, unacknowledged handover, discharge with pending result. |
| Theatre | Wrong-side attempt, checklist incomplete, implant recall trace. |
| Offline | Concurrent edits, duplicate reconnect, stale MAR, queued referral. |
| Security | Break-glass, privilege denial, expired session during medication action. |
| Human factors | Interruptions, shift change, high workload, ambiguous names, accessibility. |

# 41. Safety Sign-Off Record

| Approval | Question |
| --- | --- |
| Clinical Safety | Are identified hazards adequately controlled and evidenced? |
| Clinical Leadership | Are workflows clinically valid for intended deployment? |
| Nursing/Pharmacy/Diagnostics | Are discipline-specific safety controls acceptable? |
| Engineering | Are implemented controls consistent with approved design? |
| QA | Has independent verification passed? |
| Security/Privacy | Could security/privacy weaknesses create clinical safety risk? |
| Operations | Can the service be monitored, supported, restored and run safely? |
| Deployment Facility | Are local configuration, training and downtime arrangements ready? |

# 42. Definition of Safety Complete for a Release

System definition and intended use frozen for release candidate.

Hazard log reviewed and current.

No unaccepted Extreme residual risks.

High residual risks explicitly dispositioned by authorised governance.

Safety requirements trace to implementation and evidence.

Clinical scenario validation complete.

Known limitations documented and communicated.

Deployment configuration assessed.

Rollback/downtime/incident response ready.

Safety case signed and archived with release version.

# Appendix A - Initial Hazard Register Summary

| ID | Hazard | Domain | Initial | Residual target |
| --- | --- | --- | --- | --- |
| HAZ-ID-001 | Wrong patient selection | Identity | S4/P3 | M/L |
| HAZ-ID-002 | Duplicate patient | Identity | S4/P3 | M |
| HAZ-MED-001 | Allergy missed | Medication | S5/P2 | M |
| HAZ-MED-002 | Dose/unit error | Medication | S5/P3 | M |
| HAZ-LAB-001 | Wrong specimen identity | Laboratory | S5/P2 | M |
| HAZ-LAB-003 | Critical result unacknowledged | Laboratory | S5/P3 | H |
| HAZ-IMG-002 | Wrong laterality | Imaging | S5/P2 | M |
| HAZ-REF-001 | Referral lost | Referral | S4/P3 | M |
| HAZ-REF-002 | Responsibility gap | Referral | S5/P2 | M |
| HAZ-ADT-002 | Handover omission | ADT | S4/P3 | M |
| HAZ-ADT-003 | Pending result missed after discharge | ADT | S5/P2 | M |
| HAZ-THR-001 | Wrong patient/site/laterality | Theatre | S5/P2 | M |
| HAZ-ANAT-002 | Wrong laterality from viewer | Anatomy | S5/P2 | M |
| HAZ-CDS-002 | Alert fatigue | CDS | S4/P4 | M |
| HAZ-AI-001 | Generated summary misstates fact | AI | S4/P3 | M |
| HAZ-OFF-001 | Stale offline record | Offline | S5/P3 | M |
| HAZ-INT-002 | Semantic mapping error | Integration | S5/P2 | M |
| HAZ-INF-001 | Critical record unavailable | Infrastructure | S5/P2 | M |

# Appendix B - Safety Artefact Set

| Artefact | Purpose |
| --- | --- |
| Clinical Safety Management Plan | How safety activities are governed for project/deployment. |
| Hazard Log | Living register of hazards, controls and residual risk. |
| Safety Requirements Traceability Matrix | Hazard → requirement → implementation → test → evidence. |
| Clinical Validation Report | Scenario/workflow validation by qualified users. |
| Safety Test Report | Results of safety-focused verification. |
| Known Limitations Register | Residual constraints and unsupported use. |
| Incident/CAPA Register | Post-market safety learning. |
| Release Safety Case | Argument and evidence supporting release. |
| Deployment Safety Assessment | Facility-specific configuration/readiness review. |

# Appendix C - Relationship to Subsequent Specifications

PULSE-SAFE-001 is cross-cutting. Every specialist specification must include a safety section that references this document, identifies module-specific hazards, creates additional safety requirements where needed and supplies evidence to the central hazard log. PULSE-ANAT-001, PULSE-REF-001, PULSE-LAB-001, PULSE-IMG-001, PULSE-MED-001, PULSE-ADT-001 and PULSE-THR-001 are expected to materially expand the hazard set.

# Appendix D - Clinical Governance Note

Risk scores in this baseline are engineering/clinical-safety planning classifications, not claims of measured incident probability. They must be reviewed with qualified clinical stakeholders and adapted to the intended deployment, local protocols, professional scope and applicable regulatory obligations. The product must never use this document to bypass local clinical governance.